## Contents
Gavi 6.0 CCE Minimum Floors Eligible Product Lists under UNICEF Supply Division Long-Term Arrangements (LTAs)
Section | Sheet Name | Description
1 | Eligible_Ref_Freezer_Products | List of refrigerator and freezer products with technical specifications and indicative unit costs for 2026 and beyond.
2 | Eligible_Passive_Products | List of passive cold chain products with technical specifications and indicative unit costs. The product list includes freeze-preventative and traditional cold boxes and vaccine carriers, Long-Term Passive Devices, and Transportable Powered Vaccine Storage Devices (TPVS).
3 | Temp_Monitoring_Device | List of temperature monitoring devices with technical parameters and indicative unit costs for 2026 and beyond.
4 | Voltage_Stabilizer | List of Voltage Stabilizers with technical specifications and indicative unit costs. The product list includes freeze-preventative and traditional cold boxes and vaccine carriers, Long-Term Passive Devices, and Transportable Powered Vaccine Storage Devices (TPVS).
5 | Icepack | List of ice packs waterpacks with technical parameters and indicative unit costs for 2026 and beyond.
Caveats / Notes
1. Reference: Theproducts included in this file are prepared with reference to the Gavi Cold Chain Equipment Programme and Technology Guide, and includes products that are eligible under the Gavi 6.0 Cash Budget CCE Minimum Floors.
Please access the CCE Programme and Technology Guide here.
2. This product list is prepared by Infrastructure and Energy Centre, UNICEF Supply Division. It will be updated in discussion with Gavi as current and ongoing Long-Term Arrangements (LTAs) are finalized, and as new products, technical parameters and unit costs become available.
3. Not all products available under UNICEF LTAs are eligible for Gavi’s Cold Chain Equipment (CCE) programme. Please contact Gavi for clarification on eligibility.
4. Gavi is implementing a 40% pricing threshold limit for Ice-Lined Refrigerators (ILR) and Solar Direct Drive (SDD) equipment.
5. All refrigerators and freezers include Equipment Monitoring System (EMS) Level 2 by default. EMS Level 3 is optional and should be prioritized based on country needs and programme requirements.
6. For EMS Level 3 subscriptions, please contact UNICEF Supply Division for 1-year renewal price estimates.
7. Subscription costs associated with global SIMsfor RTMDs and EMS Level 3 are based on average estimates. Actual costs will vary by country.
Version 2.1, 1 April 2026
Prepared by Infrastructure and Energy Centre, UNICEF Supply Division

## 1.Eligible_Ref_Freezer_Products
Back to Contents
1. List of Gavi 6.0 CCE Minimum Floors Eligible Products Under UNICEF LTA: On-grid and Solar Direct Drive Refrigerators and Freezers
| Product Category | Storage Volume Category (L) | Supplier | Model | Manufactured in | PQS number | Net Storage Volume (L) | Waterpack Freezing Capacity (Kg/24h) | Total Gross weight per packing unit, kg | Energy Consumption, stable running (kWh/day) | Holdover Time (hours) | UNICEF Material number as per the Long Term Arrangments (LTAs) | UNICEF Indicative Price 2026 (USD, with EMS L2) | UNICEF Indicative Average Price 2027-2030 (USD, with EMS L3) | Note
1 | Mains Powered refrigerators | ≥30 – <60L | Aucma | CFD-50 | China | E003-079e | 50 | N/A | 164 | 0.77 | 120 | LTA to establish | 2000 | 2500 | Awaiting EMS WHO prequalifications
2 | Mains Powered refrigerators | ≥30 – <60L | Godrej | GHR 60 AC | India | E003/148 | 58 | N/A | 115 | 1.922 | 85 | S0005280 | 1613 | 2113
3 | Mains Powered refrigerators | ≥30 – <60L | Haier | HBC-79 | China | Pipeline | 59 | N/A | 127 | 0.55 | 58 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
4 | Mains Powered refrigerators | ≥30 – <60L | Vestfrost | VLS 174A AC | Denmark | E003/120e | 38 | N/A | 84.7 | 0.5 | 58 | LTA to establish | 2798.4 | 3198.4 | Awaiting EMS WHO prequalifications
5 | Mains Powered refrigerators | ≥60 – <120L | Coolfinity | IceVolt 70P | Malaysia | Pipeline | 60.4 | N/A | 44 | 1.406 | 22 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
6 | Mains Powered refrigerators | ≥60 – <120L | Godrej | GHR 120 AC | India | E003/149 | 118 | N/A | 145 | 0.55 | 80 | S0005281 | 1715 | 2215
7 | Mains Powered refrigerators | ≥60 – <120L | Haier | HBC-120 | China | E003/114e | 100 | N/A | 201 | 0.4 | 129 | S0005286 | 2657 | 3057
8 | Mains Powered refrigerators | ≥60 – <120L | Meiling | YC-125EW | China | E003/142 | 95 | N/A | 162 | 0.68 | 84 | LTA to establish | 1650 | 2150 | Awaiting EMS WHO prequalifications
9 | Mains Powered refrigerators | ≥60 – <120L | Vestfrost | VLS 304A AC | Denmark | E003/110e | 98 | N/A | 132.2 | 0.69 | 55 | LTA to establish | 3048 | 3448 | Awaiting EMS WHO prequalifications
10 | Mains Powered refrigerators | ≥120L | Coolfinity | IceVolt 300P | Malaysia | E003-122e | 241 | N/A | 190 | 2.98 | 25 | S0005336 | 3816 | 4416
11 | Mains Powered refrigerators | ≥120L | Godrej | GHR 200 AC | India | E003/136e | 226.4 | N/A | 158 | 1.67 | 35 | S0005282 | 1750 | 2250
12 | Mains Powered refrigerators | ≥120L | Haier | HBC-260 | China | E003/087e | 211 | N/A | 217 | 1.47 | 62 | S0005287 | 2034 | 2434
13 | Mains Powered refrigerators | ≥120L | Vestfrost | VLS 504A AC | Denmark | E003/113e | 242 | N/A | 188.6 | 0.63 | 55 | LTA to establish | 4165.2 | 4565.2 | Awaiting EMS WHO prequalifications
14 | Mains Powered Combined refrigerators/freezers | ≥30 – <60L | Godrej | GHR 55 FF AC | India | Pipeline | 58 | 2.4 | 155 | 1.5 | 85 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
15 | Mains Powered Combined refrigerators/freezers | ≥30 – <60L | Haier | HBCD-90 | China | E003/097e | 30 | 4 | 152 | 0.66 | 64 | S0005288 | 2314 | 2714
16 | Mains Powered Combined refrigerators/freezers | ≥60 – <120L | Godrej | GHR 110 FF AC | India | Pipeline | 108 | 2.4 | 170 | 1.7 | 85 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
17 | Mains Powered Combined refrigerators/freezers | ≥60 – <120L | Haier | HBCD-160 | China | Pipeline | 100 | 3.2 | 239 | 0.75 | 142 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
18 | Mains Powered Combined refrigerators/freezers | ≥60 – <120L | Vestfrost | VLS 096A RF AC | Denmark | Pipeline | 110 | 2.4 | 190 | 1.64 | 93 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
19 | Mains Powered Combined refrigerators/freezers | ≥120L | Haier | HBCD-220 | China | Pipeline | 146 | 3.2 | 245 | 0.76 | 142 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
20 | Mains Powered Freezer | ≥90 – <150L | Aucma | DW-25W147 | China | E003/060 | 96 | 14.5 | 85 | 3.1 | 7 | LTA to establish | 615 | 1115
21 | Mains Powered Freezer | ≥90 – <150L | Haier | HBD-130 | China | Pipeline | 118 | 15.2 | 76 | 1.33 | 2 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
22 | Mains Powered Freezer | ≥90 – <150L | Vestfrost | MF 114 | Denmark | E003/024 | 98 | N/A | 77.8 | 2.24 | 3 | LTA to establish | 1400.4 | 1800.4
23 | Mains Powered Freezer | ≥150L | Aucma | DW-25W300 | China | E003/061 | 240 | 38.3 | 133 | 3.37 | 3 | LTA to establish | 705 | 1205
24 | Mains Powered Freezer | ≥150L | Godrej | GMF 200 ECO Lite | India | E003/130 | 153 | 20.91 | 80 | 4.6 | 10 | S0005283 | 944 | 1444
25 | Mains Powered Freezer | ≥150L | Haier | HBD-265 | China | E003/131 | 211 | 32.4 | 171 | 1.41 | 11 | S0005285 | 1250 | 1650
26 | Mains Powered Freezer | ≥150L | Vestfrost | MF 314 | Denmark | E003/023 | 271 | N/A | 117.2 | 4.24 | 4 | LTA to establish | 1634.4 | 2034.4
27 | SDD refrigerators | ≥30 – <60L | Aucma | CFD-50 SDD | China | E003/098e | 50 | N/A | 268 | N/A | 120 | LTA to establish | 2875 | 3375 | Awaiting EMS WHO prequalifications
28 | SDD refrigerators | ≥30 – <60L | Haier | HTC-110 | China | E003/076e | 59 | N/A | 181 | N/A | 96 | S0005290 | 3603 | 4003
29 | SDD refrigerators | ≥30 – <60L | Vestfrost | VLS 054A SDD | Denmark | E003/106e | 55.5 | N/A | 185.8 | N/A | 90 | LTA to establish | 5307.6 | 5707.6 | Awaiting EMS WHO prequalifications
30 | SDD refrigerators | ≥60 – <120L | Dulas | VC88SDD | United Kingdom | E003-059e | 88 | N/A | 295 | N/A | 78 | LTA to establish | 6072 | 6572 | Awaiting EMS WHO prequalifications
31 | SDD refrigerators | ≥60 – <120L | Haier | HTC-120 | China | E003/116e | 100 | N/A | 260 | N/A | 112 | S0005291 | 4417 | 4817
32 | SDD refrigerators | ≥60 – <120L | Vestfrost | VLS 094A SDD | Denmark | E003/107e | 92 | N/A | 212.5 | N/A | 78 | LTA to establish | 5738.4 | 6138.4 | Awaiting EMS WHO prequalifications
33 | SDD refrigerators | ≥120L | Dulas | VC200SDD | United Kingdom | E003/040e | 132 | N/A | 295 | N/A | 80 | LTA to establish | 6306 | 6806 | Awaiting EMS WHO prequalifications
34 | SDD refrigerators | ≥120L | Haier | HTC-240 | China | E003/117e | 200 | N/A | 285 | N/A | 95 | S0005292 | 4829 | 5229
35 | SDD refrigerators | ≥120L | Vestfrost | VLS 154A SDD | Denmark | E003/108e | 170 | N/A | 211.8 | N/A | 73 | LTA to establish | 6844.8 | 7244.8 | Awaiting EMS WHO prequalifications
36 | SDD Combined refrigerator/freezer | ≥30 – <60L | Aucma | TCD-100 | China | E003/129e | 48 | 2 | 316 | N/A | 132 | LTA to establish | 3100 | 3600 | Awaiting EMS WHO prequalifications
37 | SDD Combined refrigerator/freezer | ≥30 – <60L | B Medical | TCW40CSDD | Luxembourg | Pipeline | 36.5 | 1.8 | 309.5 | N/A | 100 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
38 | SDD Combined refrigerator/freezer | ≥30 – <60L | Dulas | VC60SDD | United Kingdom | E003/084e | 57 | 2.4 | 295 | N/A | 82 | LTA to establish | 6769.2 | 7269.2 | Awaiting EMS WHO prequalifications
39 | SDD Combined refrigerator/freezer | ≥30 – <60L | Haier | HTCD-90 | China | E003/074e | 37.5 | 2.4 | 227 | N/A | 115 | S0005293 | 5143 | 5543
40 | SDD Combined refrigerator/freezer | ≥30 – <60L | Vestfrost | VLS 086A RF SDD | Denmark | Pipeline | 57.625 | 2.4 | 270 | N/A | 129 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
41 | SDD Combined refrigerator/freezer | ≥60 – <120L | Dulas | VC150SDD | United Kingdom | E003/048e | 102 | 2.4 | 295 | N/A | 78 | LTA to establish | 7806 | 8306 | waiting EMS WHO prequalifications
42 | SDD Combined refrigerator/freezer | ≥60 – <120L | Haier | HTCD-160B | China | E003/135e | 100 | 2 | 289 | N/A | 139 | S0005294 | 6516 | 6916
43 | SDD Combined refrigerator/freezer | ≥60 – <120L | Vestfrost | VLS 096A RF SDD | Denmark | E003/132e | 110 | 2.4 | 300 | N/A | 115 | LTA to establish | 9333.6 | 9733.6 | waiting EMS WHO prequalifications
44 | SDD Combined refrigerator/freezer | ≥120L | Aucma | TCD-188 | China | Pipeline | 120 | 2.2 | 400 | N/A | 125 | LTA to establish |  |  | Anticipated WHO PQS-listed pipeline product
45 | SDD Waterpacks Freezer | ≥30 – <120L | Haier | HTD-40 | China | E003/086e | N/A | 2.4 | 153 | N/A | 120 | S0005289 | 2350 | 2750
46 | SDD Waterpacks Freezer | ≥30 – <120L | Vestfrost | VFS 048 SDD | Denmark | E003/099 | N/A | 1.6 | 151.13 | N/A | 115 | LTA to establish | 3723.6 | 4123.6
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## 2.Eligible_Passive_Products
Back to Contents | [10 empty] | *Future Price=Current Price×(1+r)n; Where: r = annual price escalation rate (decimal); n = number of years
2. List of Gavi 6.0 CCE Minimum Floors Eligible Products Under UNICEF LTA: Freeze-preventative and traditional cold boxes and vaccine carriers, Long Term Passive Device and Transportable Powered Vaccine Storage Devices (TPVS)
| Product Category | Storage Volume Category (L) | Supplier | Model | Manufactured in | PQS number | Net Storage Volume (L) | Weight Empty,Kg | Cold life at +43°C: | Model coolant-pack | UNICEF Material number | UNICEF Indicative Price 2022 (USD) | UNICEF Indicative anticipated average unit Price 2026-2030 (USD)*
1 | Freeze-preventive cold boxes | 20–25 L | Qingdao Leff International Trading Co Ltd | FFCB-20L | China | E004/064 | 20 | 22.1 | 72hrs 52min | 0.6 L | S0002576 | 234 | 284.43
2 | Freeze-preventive vaccine carriers | 1–2 L | AOV International | AFVC46 | India | E004/050 | 1.5 | 4.2 | 34 hrs | 0.6 L | S0002475 | 24.04 | 29.2209
3 | Freeze-preventive vaccine carriers | 1–2 L | Trimurti Plast Containers Private Ltd. | TPVC 46 LFF | India | E004/066 | 1.5 | 4.2 | 30.48 hrs | 0.6 L | S0002585 | 23.9 | 29.0508
4 | Freeze-preventive vaccine carriers | 1–2 L | Nilkamal Limited | BCVC46LFF | India | E004/058 | 1.5 | 3.76 | 33 hrs 50 min | 0.6 L | S0002571 | 27.96 | 33.9857
5 | Freeze-preventive vaccine carriers | 1–2 L | Blowkings | BK-FF-VC-1.6L | India | E004/052 | 1.6 | 3.4 | 30 hrs | 0.6 L | S0002213 | 37 | 44.974
6 | Freeze-preventive vaccine carriers | 1–2 L | Qingdao Leff International Trading Co Ltd | FFVC-1.7L | China | E004/051 | 1.7 | 4 | 33h 42min | 0.6 L | S0002216 | 34 | 41.3274
7 | Freeze-preventive vaccine carriers |  | B Medical Systems | RCW1 | Luxembourg | E004/059 | 1.04 | 4.8 | 32hrs 07min | 0.6L | S0002573 | 283.2 | 344.2332
8 | Freeze-preventive vaccine carriers | 1–2 L | GKS | FFVC-44LR | India | E004/071 | 1.7 | 3.58 | 30 hrs 22 min | 0.4L | S0002587 | 33.9 | 41.2059
9 | Freeze-preventive vaccine carriers | 1–2 L | Gobi Technologies, Inc. | FF001A Eclipse 1.8L | China | E004/070 | 1.8 | 3.44 | 36.05 hrs | 0.6 L | S0002586 | 89 | 108.1806
10 | Freeze-preventive vaccine carriers | 2–5 L | Blowkings | BK-VC-FF 2.4 L | India | E004/065 | 2.4 | 3.56 | 39.44 hrs | 0.6 L | S0002575 | 47.7 | 57.98
11 | Traditional Cold Boxes | 5–15 L | B Medical Systems | RCW 12 | Luxembourg | E004/004 | 7 | 11.7 | 114.9 hrs | 0.6L | S0002238 | 745.2 | 905.8002
12 | Traditional Cold Boxes | 5–15 L | Apex International | AICB-156L | India | E004/045 | 5.5 | 9.4 | 107.5 hrs | 0.4L | S0002349 | 86.36 | 104.9717
13 | Traditional Cold Boxes | 5–15 L | AOV International | ACB 264 SL | India | E004/023 | 6 | 12.82 | 132 hrs 18 min | 0.4L | S0002344 | 76.78 | 93.3271
14 | Traditional Cold Boxes | 5–15 L | Nilkamal Limited | RCB 264 SL | India | E004/034 | 6 | 13.25 | 106 hrs | 0.4L | S0002362 | 86.13 | 104.6921
15 | Traditional Cold Boxes | 5–15 L | Blowkings | CB-12-CF | India | E004/018 | 12 | 20.4 | 156 hrs | 0.4L | S0002357 | 85 | 103.3186
17 | Traditional Cold Boxes | 15–25 L | B Medical Systems | RCW 25 | Luxembourg | E004/005 | 20 | 15.9 | 134.6 hrs | 0.6L | S0002235 | 621.6 | 755.5628
18 | Traditional Cold Boxes | 15–25 L | Apex International | AICB 444 L | India | E004/010 | 18 | 19.9 | 140 hrs | 0.6L | S0002244 | 104.63 | 127.1791
19 | Traditional Cold Boxes | 15–25 L | AOV International | ACB 503 L | India | E004/015 | 18 | 21.8 | 126 hrs 32 min | 0.3L | S0002269 | 106.9 | 129.9383
20 | Traditional Cold Boxes | 15–25 L | Blowkings | CB-20-CF | India | E004/025 | 20 | 22 | 138 hrs | 0.4L | S0002232 | 120 | 145.8615
21 | Traditional Cold Boxes | 15–25 L | Nilkamal Limited | RCB 444L-A | India | E004/036 | 20.25 | 20.525 | 152 hrs 28min | 0.4L | S0002363 | 101 | 122.7668
22 | Traditional Cold Boxes | 15–25 L | Apex International | AICB 503 L | India | E004/031 | 22.5 | 19.73 | 128 hrs 16 min | 0.3L | S0002350 | 104.63 | 127.1791
23 | Traditional Cold Boxes | 15–25 L | Qingdao Leff International Trading Co Ltd | FHCB23-0624 | China | E004/069 | 23 | 18.66 | 134 hrs 39 min | 0.6L | S0002584 | 177 | 215.1458
24 | Traditional Cold Boxes | 15–25 L | Nilkamal Limited | RCB503L | India | E004/067 | 23.7 | 19.15 | 131 hrs and 5 min | 0.3L | S0002583 | 99 | 120.3358
25 | Traditional vaccine carriers | 1–2 L | AOV International | AVC 44 | India | E004/008 | 1.35 | 2.2 | 40 hrs 02 min | 0.4 L | S0002348 | 9.29 | 11.2921
26 | Traditional vaccine carriers | 1–2 L | Apex International | AIVC-44LR | India | E004/044 | 1.67 | 1.99 | 35.7 hrs | 0.4 L | S0002366 | 10.75 | 13.0668
27 | Traditional vaccine carriers | 1–2 L | Blowkings | BK - VC 1.7 - CF | India | E004/021 | 1.7 | 1.6 | 38 hrs | 0.4 L | S0002204 | 8.9 | 10.8181
28 | Traditional vaccine carriers | 1–2 L | Nilkamal Limited | BDVC44 | India | E004/054 | 1.7 | 1.65 | 38 hrs 37 min | BIP-4 | S0002574 | 9.15 | 11.1219
29 | Traditional vaccine carriers | 2–5 L | B Medical Systems | RCW 4 | Luxembourg | E004/002 | 3 | 3.1 | 30.3 hrs | 0.6L | S0002353 | 397.2 | 482.8017
30 | Traditional vaccine carriers | 2–5 L | AOV International | AVC 46 | India | E004/009 | 2.5 | 2.98 | 50 hrs 12 min | 0.6L | S0002353 | 11.52 | 14.0027
31 | Traditional vaccine carriers | 2–5 L | Blowkings | BK - VC 2.6 - CF | India | E004/020 | 2.6 | 1.9 | 43 hrs | 0.6L | S0002224 | 14 | 17.0172
32 | Traditional vaccine carriers | 2–5 L | Giostyle SpA | Gio'Style VC 2.6 L | Italy | E004/032 | 2.6 | 1.8 | 41 hrs 29 min | 0.4L | S0002211 | 35.088 | 42.6499
33 | Traditional vaccine carriers | 2–5 L | Nilkamal Limited | BCVC46 | India | E004/040 | 2.7 | 2.32 | 46 hrs 23 min | 0.6L | S0002361 | 12.62 | 15.3398
34 | Traditional vaccine carriers | 2–5 L | Apex International | AIVC-46 | India | E004/047 | 2.9 | 2.63 | 51 hrs 8 min | 0.6L | S0002351 | 14.45 | 17.5642
35 | Traditional vaccine carriers | 2–5 L | Blowkings | BK-VC 3.4 -CF | India | E004/043 | 3.4 | 2.26 | 43.43 hrs | 0.6L | S0002358 | 16 | 19.4482
36 | Long-term passive devices ( Extra long range) | 5-15 L | Qingdao Aucma Global Medical Co.,Ltd. | ARKTEK-YBC-5 | China | E004/041 | 5.4 | 22 | 35 days | 1L | S0002233 | 2393 | 2908.7222
37 | Transportable Powered Vaccine Storage Devices (TPVS) | 1–2 L | BlackFrog Technologies Private Limited | Emvolio Plus | India | E003-134 | 1.55 | 11 | 12 hrs |  | To establish | 2500 | 3038.7821
38 | Transportable Powered Vaccine Storage Devices (TPVS) | 1–2 L | MK TEC GmbH | Indigo 201 | Germany | E003/140 | 1.7 | 11 | 120 hrs |  | To establish | 2073.75 | 2520.6697
39 | Transportable Powered Vaccine Storage Devices (TPVS) | Charger | MK TEC GmbH | Indigo 201 | Germany | E003/140 | N/A | 34 | N/A |  | To establish | 3045 | 3701.2365
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## 3.Temp_Monitoring_Device
Back to Contents
3. List of Gavi 6.0 CCE Minimum Floors Eligible Products Under UNICEF LTA: Temperature Monitoring Device
Column1 | Product Category | Description | Manufacturer | Model | PQS Code | Manufactured in | Sensor Minimum & Maximum temperature measurement: | UNICEF Material number | Activated life: | 2026: Hardware+First Year subscription [USD/unit] | 1 year renewal price (G) | UNICEF Indicative Average Price 2027-2030 Hardware+First Year subscription [USD/unit]
1 | Remote Temperature Monitoring Device | Application 1: Configured for monitoring one Walk-in Cold Room (WICR) | Beyond Wireless | ICE3 - BC141 | E006/036 | South Africa | 4 x wired sensors with measurement range -40°C to +100°C | S0002548 | 7 Years | 865 | 365 | 0
2 | Remote Temperature Monitoring Device | Application 1: Configured for monitoring one Walk-in Cold Room (WICR) | Myfridgeonline | VM1000 | E006/061 | South Africa | 4 x wired sensors with measurement range -30°C to +55°C | S0002551 | 5 Years | 641.08 | 290.08 | 0
3 | Remote Temperature Monitoring Device | Application 1: Configured for monitoring one Walk-in Cold Room (WICR) | Haier | U-Cool LoRa | E006/075 | China | 5 x wireless sensors with measurement range -40°C to +80°C (can be expanded up to 60 sensors) | S0002550 | 7 Years | 1206 | 161 | 0
4 | Remote Temperature Monitoring Device | Application 1: Configured for monitoring one Walk-in Cold Room (WICR) | Nexleaf | CTX | E006/091 | India | 4 x wireless sensors with measurement range -30°C to +55°C (can be expanded up to 16 sensors) | S0002552 | 5 Years | 643 | 240 | 0
5 | Remote Temperature Monitoring Device | Application 2: Configured for monitoring two Walk-in Cold Rooms (WICRs) or one Combo Walk-in Cold and Freezer Room (WIC/FR) | Beyond Wireless | ICE3 Extra - BC440 | E006/037 | South Africa | 8 x wired sensors with measurement range -40°C to +100°C and 1 x wired sensor with measurement range -200°C to +150°C (can be expanded up to 16 sensors) | S0002554 | 3 Years | 1525 | 425 | 0
6 | Remote Temperature Monitoring Device | Application 2: Configured for monitoring two Walk-in Cold Rooms (WICRs) or one Combo Walk-in Cold and Freezer Room (WIC/FR) | Myfridgeonline | VM1000R | E006/061 | South Africa | 4 x wired sensors with measurement range -30°C to +55°C and 4 x wireless sensors with measurement range -30°C to +55°C | S0002557 | 5 Years | 1141.08 | 290.08 | 0
7 | Remote Temperature Monitoring Device | Application 2: Configured for monitoring two Walk-in Cold Rooms (WICRs) or one Combo Walk-in Cold and Freezer Room (WIC/FR) | Haier | U-Cool LoRa | E006/075 | China | 9 x wireless sensors with measurement range -40°C to +80°C (can be expanded up to 60 sensors) | S0002556 | 7 Years | 1968 | 308 | 0
8 | Remote Temperature Monitoring Device | Application 2: Configured for monitoring two Walk-in Cold Rooms (WICRs) or one Combo Walk-in Cold and Freezer Room (WIC/FR) | Nexleaf | CTX | E006/091 | India | 8 x wireless sensors with measurement range -30°C to +55°C (can be expanded up to 16 sensors) | S0002558 | 5 Years | 842 | 240 | 0
9 | Remote Temperature Monitoring Device | Application 3: Configured for monitoring three Walk-in Cold Rooms (WICRs) | Beyond Wireless | ICE3 Extra - BC440 | E006/037 | South Africa | 8 x wired sensors with measurement range -40°C to +100°C and 5 x wired sensors with measurement range -200°C to +150°C (can be expanded up to 16 sensors) | S0002560 | 3 Years | 1835 | 485 | 0
10 | Remote Temperature Monitoring Device | Application 3: Configured for monitoring three Walk-in Cold Rooms (WICRs) | Haier | U-Cool LoRa | E006/075 | China | 13 x wireless sensors with measurement range -40°C to +80°C (can be expanded up to 60 sensors) | S0002562 | 7 Years | 2740 | 455 | 0
11 | Remote Temperature Monitoring Device | Application 3: Configured for monitoring three Walk-in Cold Rooms (WICRs) | Nexleaf | CTX | E006/091 | India | 12 x wireless sensors with measurement range -30°C to +55°C (can be expanded up to 16 sensors) | S0002563 | 5 Years | 1036 | 240 | 0
12 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | Berlinger | SITE L | E006/087 | Sri Lanka | 1 x wired sensor with measurement range -40°C to +75°C | S0002566 | 3 Years | 415 | 70 | 0
13 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | eHealth | Trek Pro & Passport | E006/080 | Thailand | 1 x wireless sensor with measurement range -30°C to +70°C (can be expanded up to 200 sensors total) | S0002581 | 3 Years | 455 | 108 | 0
14 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | Beyond Wireless | ICE3 - BC141 | E006/036 | South Africa | 1 x wired sensor with measurement range -40°C to +100°C (can be expanded up to 4 sensors) | S0002565 | 7 Years | 570 | 320 | 0
15 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | Haier | U-Cool | E006/060 | China | 1 x wired sensor with measurement range -40°C to +120°C | S0002567 | 5 Years | 272 | 161 | 0
16 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | Myfridgeonline | VM1000 | E006/061 | South Africa | 1 x wired sensor with measurement range -30°C to +55°C (can be expanded up to 4 sensors) | S0002568 | 5 Years | 562.08 | 290.08 | 0
17 | Remote Temperature Monitoring Device | RTMD-(Refrigerator) | Nexleaf | CTX | E006/091 | India | 1 x wireless sensor with measurement range -30°C to +55°C (can be expanded up to 16 sensors) | S0002570 | 5 Years | 329 | 80 | 0
18 | 30-day electronic temperature logger | 30DTR, external sensor (For Ref.) | Berlinger | Fridge-tag 2 E, ext sensor | E006/040 | Switzerland/Malaysia | -40°C to +60°C | S0002135 | 60 months | 90.3 | N/A | 0
19 | 30-day electronic temperature logger | 30DTR, external sensor (For Frz.) | Berlinger | Fridge-tag 2 E, ext sensor | E006/040 | Switzerland/Malaysia | -40°C to +60°C | S0002179 | 60 months | 90.3 | N/A | 0
20 | 30-day electronic temperature logger | 30DTR, internal sensor (For Ref.) | Berlinger | Fridge-tag 2 E, int sensor | E006/040 | Switzerland/Malaysia | -25°C to +655C | S0002136 | 60 months | 68.3 | N/A | 0
21 | 30-day electronic temperature logger | 30DTR, internal sensor (For Ref.) | Haier | HETL-01 | E006/042 | China | -20°C to +50°C | S0002128 | 24 months | 23 | N/A | 0
22 | 30-day electronic temperature logger | 30 DTR, internal sensor (For Ref.) | LogTag North America Inc. | VaxTag+ | E006/099 | Vietnam | -30°C to +65°C | S0005273 | 48 months | 37.9 | N/A | 0
23 | 30-day electronic temperature logger | 31 DTR, internal sensor (For Frz.) | LogTag North America Inc. | VaxTag+ | E006/099 | Vietnam | -40°C to +60°C | S0005274 | 45 months | 38.9 | N/A | 0
24 | 30-day electronic temperature logger | 30DTR, internal sensor (For Ref.) | G-Tek | LM Pro FZT | E006/101 | India | 30°C to +60°C | S0005272 | 60 months | 65 | N/A | 0
25 | Irreversible Freeze Indicator | Freeze Indicator | Berlinger | Freeze-tag® | E006/007 | Malaysia | -20°C to +55°C | S0002129 | 60 months | 8.65 | N/A | 0
26 | Irreversible Freeze Indicator | Freeze Indicator | Sensitech Inc. | FreezeAlert | E006/009 | The Netherlands | -20°C to +55°C | S0002134 | 36 months | 3.55 | N/A | 0
27 | Irreversible Freeze Indicator | Freeze Indicator | LogTag | LogTag ISOTag | E006/038 | China | -25°C to +60°C | S0002194 | 36 months | 9.6 | N/A | 0
28 | User programmable temperature data logger | Data logger | LogTag | Utrid-16F | E006/076 | China | -20°C to +55°C | S0002291 | 24 months | 27.2 | N/A | 0
29 | User programmable temperature data logger | Data logger | Parsyl | Trek Pro (TrekA2) | E006/069 | China | -25°C to +60°C | S0002292 | 36 months | 127 | N/A | 0
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## 4.Voltage_Stabilizer
Back to Contents
4. List of Gavi 6.0 CCE Minimum Floors Eligible Products Under UNICEF LTA: Temperature Monitoring Device,Voltage Stabilizer & Ice Packs
| Product Category | Description | Manufacturer | Model | PQS Code | Manufactured in | Sensor Minimum & Maximum temperature measurement: | UNICEF Material number | Activated life: | UNICEF Indicative Price 2026 (USD) | UNICEF Indicative anticipated average unit Price 2026-2030 (USD)* | Note
1 | Voltage Stablizers | 1kVA 230V 1ph ExtRange | Haier | HVS-1000U | E007/009 | China | 110-264V/50-60 Hz | TBD | 24 months | 250 | 0
2 | Voltage Stablizers | 1kVA 230V 1ph ExtRange | Sollatek | SVS500-22E5 | XXXX | TBD | 110-264V/50-60 Hz | TBD | 24 months | to be confirmed | to be confirmed | Anticipated WHO PQS-listed pipeline product

## 5.Icepack
Back to Contents
5. List of Gavi 6.0 CCE Minimum Floors Eligible Products Under UNICEF LTA: Ice packs and Waterpack
| Product Category | Description | Manufacturer | Model | PQS Code | Manufactured in | UNICEF Material number | UNICEF Indicative Price 2022 (USD) | UNICEF Indicative anticipated average unit Price 2026-2030 (USD)*
1 | Icepack | Icepack,0.3 litre capacity | AOV International | AIP-3 | E005/004 | India | S0002095 | 0.26 | 0
2 | Icepack | Icepack,0.3 litre capacity | Blowkings | BK-V4-H | E005/010 | India | S0002095 | 0.3 | 0
3 | Icepack | Icepack,0.3 litre capacity | Nilkamal Limited | BIP-3 | E005/013 | India | S0002095 | 0.27 | 0
4 | Icepack | Icepack,0.3 litre capacity | Apex International | AIIP-03 | E005/007 | India | S0002095 | 0.38 | 0
5 | Icepack | Icepack,0.3 litre capacity | Qingdao Leff International Trading Co Ltd | WP-0.3L | E005/020 | China | S0002095 | 0.27 | 0
6 | Icepack | Icepack,0.4 litre capacity | AOV International | AIP-4 | E005/005 | India | S0002220 | 0.26 | 0
7 | Icepack | Icepack,0.4 litre capacity | Apex International | AIIP-04 | E005/008 | India | S0002220 | 0.4 | 0
8 | Icepack | Icepack,0.4 litre capacity | Blowkings | BK-4 | E005/011 | India | S0002220 | 0.3 | 0
9 | Icepack | Icepack,0.4 litre capacity | Nilkamal Limited | BIP-4 | E005/014 | India | S0002220 | 0.28 | 0
10 | Icepack | Icepack,0.4 litre capacity | Giostyle SpA | Water-pack 0.4 L | E005/019 | Italy | S0002220 | 0.7 | 0
11 | Icepack | Icepack,0.4 litre capacity | Qingdao Leff International Trading Co Ltd | WP-0.4L | E005/022 | China | S0002220 | 0.27 | 0
12 | Icepack | Icepack,0.6 litre capacity | AOV International | AIP-6 | E005/006 | India | S0002221 | 0.35 | 0
13 | Icepack | Icepack,0.6 litre capacity | Blowkings | BK-6 | E005/012 | India | S0002221 | 0.45 | 0
14 | Icepack | Icepack,0.6 litre capacity | Nilkamal Limited | BIP-6 | E005/015 | India | S0002221 | 0.42 | 0
15 | Icepack | Icepack,0.6 litre capacity | Apex International | AIIP-06 | E005/009 | India | S0002221 | 0.48 | 0
16 | Icepack | Icepack,0.6 litre capacity | Qingdao Leff International Trading Co Ltd | WP-0.6L | E005/021 | China | S0002221 | 0.35 | 0
17 | Waterpack | 1L Waterpack for Arktek-YBC-5 | Qingdao Aucma Global Medical Co.,Ltd. | Waterpack Arktek-YBC-5 1L | E004/041 | China | S0002297 | 9 | 0
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.